## Table 1
TECHNICAL SHEET No.:
Equipment Name: 2D LAPAROSCOPY TOWER FOR GENERAL SURGERY
No. | Configuration Kits | Edge | Yes | Not | Obs.
Laparoscopy Tower 2d For General Surgery
1 | Camera Control Unit 2d 2d Camera Platform That Can Adapt to Different 2d Camera Heads The Unit Allows Camera Heads To Be Connected / Disconnected During Surgical Procedures25 To Have At Least 5 Preinstalled Standard Profiles (General Surgery Lap, Urology Lap, Gynecology Lap, Neurosurgery, Cardio-Thoracic Surgery) To have the possibility to adjust settings for each profile - Brightness (-5 To +5, In Steps Of 1) - Digital Zoom (1.0, 1.2, 1.4, 1.6 and 1.8) - Contrast (-5 to +5, in steps of 1) - Peripheral Image Improvement (-5 to +5, in steps of 1) to have the possibility of adding at least 4 effects: Red Color Intensification Effect For Visualizing Different Structures - Smoke Reduction Effect On Display - Pop (Picture Out Of Picture) To Allow Two Images To Be Displayed Simultaneously (Ex: Standard Image Vs Image With Red Color Intensification) - On-Screen Indicators (Image Status) - The Red Color Intensification Effect Can Be Used Simultaneously With The Image Enhancement Algorithm In Smoke Conditions The Camera Menu Allows Saving At Least 10 Favorite Functions With Quick Access The camera menu allows saving at least 3 camera settings profiles, along with the pre-installed ones The Current Settings Of The Camera Can Be Displayed On The Screen In The Form Of Text Or Symbol The unit is compatible with the control functions of the video endoscopic camera (menu, white calibration, digital zoom, turning on/off the light source, starting and stopping the video recording /Snapshot)The buttons of the video endoscopic camera / on the unit can be programmed with at least 8 functions according to the needs of the specialist To have a white balance function on the unit and on the video endoscopic camera To allow the automatic control of the light intensity in combination with the LED light source, through the Mis-Bus connection To have a gain function, for automatically increasing the sensitivity to light, permanently active The menu must be available in at least 9 languages Minimum Connections Required: - 2 X Dvi-D (1080p) - 2 X 3g-Sdi (1080p) - 1 X HD-Sdi (1080i) - Earthing connector - Power connector - Mis-Bus connector - 2 X 3.5 mm Jack - 1 X Lan Ethernet, Isolated (For Service) - 1 X Usb | 1
2 | Head Video Camera 2d 1 Cmos With Optical Zoom Sensor Resolution: Minimum Full Hd Native (1920 X 1080 Pixels) | 1
To have a 1/3" Cmos Sensor to have Progressive ScanTo have a focal length of 14 to 28 mm to have a shutter: 1/6 to 1/60,000 S The camera can be disinfected by wiping or immersion The camera can be cleaned with alkaline and thermal solutions The camera can be connected to the 2d and 3d platform To have at least 4 programmable buttons that can be configured to control the display, photo/video recordings, menu navigation and change camera settings The functions of the buttons on the camera head should correspond to the functions of the buttons on the unit The buttons can be programmed with at least 8 functions according to the user's requirements To Have Preset Functions on the Camera Head in Live Mode: Menu, White Balance, Digital Zoom, Light Source On/Off, Photo/Video Recording To Allow Automatic Light Adjustment
3 | Telescope For Laparoscopy 2d Full Hd Telescope For Minimally Invasive Laparoscopic Interventions To Be Compatible With 2d CameraTo have a length of 330 mm to have a diameter of 10 mm To have a viewing angle of 30 degrees To be autoclavable at 134° C | 1
4 | Telescope For Laparoscopy 2d Full Hd Telescope For Minimally Invasive Laparoscopic Interventions To Be Compatible With 2d Camera To have a length of 330 mm to have a diameter of 10 mm To have a viewing angle of 0 degrees To be autoclavable at 134° C | 1
5 | Sterile Covers For 2d Camera Head, Ring Type, Compatible With Camera Head, Single Use | 50
6 | Basket for Sterilizing 2 Laparoscopes with a Maximum Length of 340mm, Outer Size 454x104x41mm | 1
7 | Medical Monitor 2d Full Hd 27" Be Certified For Medical Use To have a TFT display with IPS technology To have LED lighting with ultra strong luminance to have a diagonal of at least 27" To have Aspect 16:9 Monitor resolution: Min. 1920 X 1080 Pixels (Full Hd) To Have a Liquid-Resistant ScreenTo have a menu in at least 8 languages Navigation through the menu should be done through OSD (On Screen Display) Be Able to Change Language, Source, Saturation, Tint, Contrast, Brightness, Color, Gamma, Video Level, Color Corrector, Auto-Select Source, Menu Lock Have at least the following Color Temperature Levels: 5500, 6500, 8000, 9300, "User" To Have Pincture-In-Picture Function With the Possibility of Controlling the Image Size (Small, Large, Divided, Inverted) Minimum Video Entries: - 1 X Dvi-I - 1 X Dvi-D - 2 X 3g-Sdi - 1 X VGA-Rgbs / Ypbpr-Sog Minimum Video Outputs: Dvi-I, Dvi-D, 2 X 3g-Sdi - 1 X Dvi-I - 1 X Dvi-D - 2 X 3g-Sdi | 1
Pixel Size: Minimum 0.311mm Display Size: Minimum 598mm Horizontal X 336mm Vertical Number of Colors: Minimum 1.07 Billion Minimum Contrast: 1000:1 The maximum brightness should be at least 850 Cd/M2 at 6500kHorizontal/Vertical Viewing Angle Must Be Minimum 178° Vesa 100 Standard
8 | Medical Touchscreen Monitor To have a certificate for medical use To have a display with TFT technology and LED lighting to have a diagonal of at least 21.5 inches To have an Lcd panel with a protective layer against reflections. To have a panel with a touch function To have Aspect 16:9 Resolution: At least 1920 X 1080 Pixels (Full Hd) to have a menu in at least 6 languages To Allow Menu Navigation With On Screen Display To Have the Possibility of Changing Language, Source, Contrast, Brightness, Color Mode, Color) Allow Color Adjustment: Red, Green, Blue Have Color Modes: Cold, Neutral, Warm, Customized Have the Possibility of Reset to Factory Settings Minimum Necessary Inputs: Dvi-D, VGA, Hdmi, Display Port, Usb Pixel Dimensions: Minimum 0.2475 Mm Horizontal X 0.2475 Mm Vertical To Have Minimum 16.7 Billion Colors To have Contrast 1000:1To have a maximum brightness of at least 250 Cd/M2 To have a horizontal/vertical viewing angle of at least 178° | 1
9 | Cool Led Light Source To Have Front Panel With Lcd Display Average Life Time For Led Module To Be At Least 30,000 Working Hours The display must be monochrome with a diagonal of at least 2.4" / 6 cm to indicate the operation status (Standy), light level, automatic light control, connection with the light cable, light cable test, temperature, color, LED module, to be approx. 5665 K (±6.3%) Luminous Flux Of The Led Module To Be 2,510 Lumen (±3.6%) Light Intensity Control Can Be Made With The Rotary Handle With 5% Steps (0-100%) To Have Laser Mode For Infrared Light Applications: The Image Agent To Be Indocyanine Green The laser must be Class 1 M according to IEC 60825-1. The laser center wavelength must be 805 Nm (+5/-10). The spectral width of the laser must be 3 Nm.Automatic Light Intensity Control Can Be Made Through the Room Control Unit (Misbus) To allow deactivation of the automatic lighting mode by manual control of the light intensity with the help of the rotary handle To Have a Bus-Type Control Interface (Source Start / Standby) Through the Connected Camera Head The connector for the light cable should be adaptable to several manufacturers, for Aesculap/Storz, Olympus and Wolf light cables with diameters of 3.5 / 4.8 mm To Have Integrated Anti-Glare Protection (In The Case In Which The Light Cable Is Not Connected To The Unit) To Automatically Turn Off The Light When The Light Cable Is Removed From The Unit To have the Function of Controlling the Functionality of the Cables Light with display on the LCD screen (for cables with a diameter of 4.8 mm) | 1
10 | Full Hd Optical Cable for Minimally Invasive Laparoscopic InterventionsTo be compatible with Aesculap/Storz/Olympus telescopes | 1
To be compatible with the light source and to have a length of 2.5 meters To have a minimum diameter of 4.8 mm To be resistant to sterilization by autoclaving at 134 degrees To Have a Mechanical Anti-Disengagement Coupling System to the Telescope, What It allows the axial rotation of the optical cable in relation to the connector on the optical telescope without coming apart
11 | Fiber Optic Adapter To be compatible with the light source to be reusable To be compatible with optical fibers Aesculap / Storz / Olympus / Wolf/Aesculapius | 1
12 | Co2 inspirator 50l/Min To Have Front Panel With Lcd To have a color display of at least 5.7" with touchscreen to have the maximum gas flow: Min. 50 L/Min To Have Maximum Pressure: Minimum 75 Mmhg To Have a Pressure Adjustment Range Between 1-30 Mmhg To have the Inlet Pressure of Min. 2.7 Bars (For Hospital Co2 Station) And Maximum 80 BarsTo have a smoke evacuation flow rate of 6 L/Min or 12 L/Min To have at least 4 modes: Standard, Bariatric, Pediatric, Cardio (For Endoscopic Assisted Cardiac Surgery) To have an integrated gas heating function To have a universal door buzzer for connecting to the hospital or cylinder CO2 station To have a display with playback in at least 17 languages To have at least 5 levels of gas supply monitoring To display the gas used in liters To Monitor Gas Supply Through Symbols And Acoustic Signals | 1
13 | Inflator Connector Tube to Co2 Cylinder To be compatible with the inflator To be at least 1.5 meters long To have a connector for a door-type device to have a connector for a Din-type cylinder To withstand pressures of up to 80 bars | 1
14 | Carbon Case for Smoke Evacuation To be compatible with the inflatorTo contain active carbon for gas filtration, to filter gases and inhaled smoke To have a hydrophobic filter | 30
15 | Piping For Smoke Evacuation To be compatible with the inflator to contain a hydrophobic filter To prevent contamination of the device in case of accidental aspiration of fluids To have a length of at least 3 m and to be of single use To be Sterile Packed | 50
16 | Filter For Co2 Inspirator To be compatible with the inflator To filter the carbon dioxide breathed into the patient, to have a hydrophobic antibacterial filter To Be Single Use To have an efficiency of at least 99.9999% to have a 22mm Iso type connector To be Sterile Packed | 50
17 | Co2 insufflator tubing To be compatible with the inflator to heat the gas being blown To be resistant to at least 100 sterilizations by autoclaving at 134 Of Degrees | 2
To be made of silicone To have a Luer-Lock type connector
18 | Irrigation and Suction PumpMultifunctional System For At Least 4 Specialties: Laparoscopy, Arthroscopy, Hysteroscopy, Urology With Different Color Codes To have the possibility for each mode to be activated with a transponder with a specific RFID, only needed once System With Vacuum For Suction And Peristaltic Pump For Irrigation To Have Color Display With Touchscreen To allow the use of reusable and single-use tubes with automatic identification chip for each specialization Each Tubular With Chip Allows The Activation Of Specific Settings As A Function Of Specialization Allow Activation From: - Infrared remote control - Pedal - Instrument - Directly From the Unit Laparoscopy Mode - Have Operating Modes: High Flow (Flow Up to 3.5 L/Min), Standard (Flow Between 1 L/Min and 2.5 L/Min) - To have a Maximum Irrigation Flow of 3.5 L/Min Arthroscopy Mode To Have As Operating Modes: Knee, Umar, Hip, Small Joints - Each mode must have a pressure and a flow set from the factory that can be changed later: For small joints 35 Mmhg / 0.7 L/Min, For knees 45 Mmhg / 1.0 L/Min, For Umar 50 Mmhg / 1.5 L/Min, For Balance 60 Mmhg / 1.0 L/Min - To allow the setting of 2 different users - Have Maximum Irrigation Flow: 2.5 L/Min - To have the pressure between 15-200 Mmhg - To Permanently Display the Current Pressure Hysteroscopy Mode - To have a maximum irrigation flow of 0.5 L/Min - To have pressure between 15-150 Mmhg Urology Module - To have a maximum irrigation flow of 0.5 L/Min - To have a pressure between 15-150 mmhg and to have a menu in at least 5 languages To have an audible warning for overpressure. To have an audible warning for a contaminated tubeTo have the maximum negative pressure for suction of 700 Mbar to have the maximum output pressure of 500 Mmhg ±10% Minimum connections required: - Service interface - Pedal connection - Food Connection | 1
19 | Laparoscopy Transponder For Irrigation/Aspiration Unit | 1
20 | Irrigation / Suction Tubing For Laparoscopy Mode To Be Single Use To have a minimum length of 450 cm To have Multi-Adapter for Suction Connection To be equipped with a faucet operated with buttons for aspiration and irrigation To Have Connections With Those | 50
21 | Single Use Suction / Irrigation Cannula The cannula for irrigation and aspiration must be at least 330 mm long To be provided with holes along the tip to have a diameter of 5 mm To be Sterile Packed | 50
22 | Suction Jar, Capacity 2l, Single Use | 40
23 | Vacuum Tube Connecting Suction Unit To Jar, With Hydrophobic Filter | 10
24 | Unit For Vascular SealingTo be controlled by a microprocessor To Have Self-Test Function When Starting the Unit To have a permanent self-test function, including during use To Allow Automatic Detection When Connecting the Instrument To Allow Control Through the Pedal and Button of the Instruments To have at least two modes of use: Standard - Plus (With Increased Power) - Acoustic Alarm During the Sealing Process Have a Display for the Status of the Sealing Process (Ready for Sealing, Sealing in Progress, Sealing Completed) To Display Errors With Unique Code With Acoustic Alarm | 1
25 | Bipolar Vascular Sealing Forceps, Non-articulated Jaws, Jaw Length 26.5mm, Cutting Blade Included, Cutting Length 23.5mm, Diameter 5mm, Length 360mm, Sterile, Single Use | 6
26 | Bipolar Vascular Sealing Forceps, Articulated Jaws, Jaw Length 26.5mm, Cutting Blade Included, Cutting Length 23.5mm, Diameter 5mm, Length 360mm, Sterile, Single Use | 6
27 | Bipolar Vascular Sealing Forceps, Maryland Type Jaws, Non-articulated, Jaw Length 21.5mm, Cutting Blade Included, Cutting Length 20mm, Diameter 5mm, Length 360mm, Sterile, Single Use | 6
28 | Bipolar Vascular Sealing Forceps, Maryland Type Jaws, Articulated, Jaw Length 21.5mm, Cutting Blade Included, Cutting Length 20mm, Diameter 5mm, Length 360mm, Sterile, Single Use | 6
29 | Endoscopy winch To be a modular system that can be adapted for several departments to have a "soft close" drawer with a key To have at least 4 shelves, of which 3 with height adjustment and 1 with handles for handling To have 4 anti-static wheels with braking system, (diameter of 125 mm) To have channels for integrated cables in the side columns To have at least 2 panels in the back for easy access to the cables To have a support for the transformer and to have a side button for power supply To be compatible with a wide range of accessoriesTotal Dimensions Without Accessories: Max. 703 X 1506 X 663 Mm Shelf Dimensions: Min. 450 X 485 Mm To have at least 12 power outlets for the units to have Weight without Transformer: Max. 55 Kg Weight With Transformer: Max. 75 Kg To Support Maximum Weight Per Shelf: Min. 40 Kg To Have The Maximum Allowed Weight Of Min. 210 Kg To have a transformer | 1
30 | 2d Camera Head Support, Dimensions Approx. 32 X 35 mm | 1
31 | Support for Co2 cylinder, for cylinders between 80 and 200 mm in diameter, capable of supporting at least 39 kg, compatible with the winch | 1
32 | Standard Bar For Endoscopy Winch, Capable Of Supporting At Least 9 Kg, Compatible With The Winch | 1
33 | Basket For Accessories, Dimensions Approx. 300 X 300 X 200mm, Able To Support Minimum 4 Kg, Compatible With Winch | 1
34 | Additional shelf with location on the side of the winch, dimensions approx. 330 X 330 Mm, Able To Support Weight Of Minimum 14 Kg | 1
35 | Back Panel For Endoscopy Winch | 1
36 | Support For Fixing The Winch Power Cable, With Holes For Plug Fixing | 1
37 | General Power Cable Winch To have a minimum length of 5 m and to have a Euro type plug | 1
38 | Winch Power Transformer | 1
To have a maximum power consumption of 2200 Va To be equipped with a 230v monitoring relay and to have protection against Ip20 type liquids
39 | Arm For Central Monitor, Adjustable At Minimum Height ±160 Mm, Minimum Length 40 Cm Capable Of Supporting Monitor Weight Between 7 - 12 Kg | 1
40 | Crossbar For Winch Reinforcement, For Fixing Monitor Support Arm, Compatible With Winch | 1
41 | Arm for Side Monitor, With Extension of Minimum 45 Cm, Height Adjustable ±335 Mm, Minimum Arm Length 75 Cm, Capable of Supporting Monitor Weight Between 5 and 8.5 Kg | 1
42 | Earthing Cable, Minimum Length 1.5 M | 2
43 | Earthing Cable, Minimum Length 4 M | 7
44 | Power Cable Extension, Minimum Length 1 M | 6
45 | Power Cable Extension, Minimum Length 2.5 M | 2
46 | Clamp For Securing Cables, Dimensions Approx. 138 X 20 Mm, Compatible With The Winch | 1
Basic Laparoscopic Instrument Kit
47 | Veress Insufflation Needle, Diameter 2.1mm, Length 150mm | 1
48 | Trocar Body, Smooth, With Faucet, Diameter 5mm, Length 110mm, Red Color | 3
49 | Triangular Obturator For Trocar, Diameter 5mm, Length 110mm, Color Red | 3
50 | Complete Silicone Gasket For 5mm Swap, Red Color | 20
51 | Trocar Body, Smooth, With Faucet, Diameter 10mm, Length 110mm, Color Green | 4
52 | Triangular Obturator For Trocar, Diameter 10mm, Length 110mm, Green Color | 4
53 | Full Silicone Gasket For 10/12mm Swap, With 5mm Rocker Valve Type Reduction | 3
54 | Full Silicone Gasket For 10/12mm Swap | 1
55 | Trocar Body, Smooth, With Faucet, Diameter 12mm, Length 110mm, Color Yellow | 1
56 | Triangular Obturator For Trocar, Diameter 12mm, Length 110mm, Color Yellow | 1
57 | Full Silicone Gasket For 10/12mm Swap | 1
58 | Metzenbaum Monopolar Scissors, Black, Non-Reflective, Tungsten Inserts, Curved, Butt/Butt Tip, Fine Serrated, Rotary, Insulated, Double Action, Complete Standard Tool, Detachable, Composed of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert Sheath And Insert), Diameter 5mm, Length 310mm | 2
59 | Hook Type Monopolar Scissors, Straight, Butt/Butt Tip, Rotary, Insulated, Simple Action, Complete Standard Tool, Detachable, Composed of 4 Distinct Modular Elements (Handle, Insulating Sheath, Sheath Insert and Insert), Diameter 5mm, Length of Active Part 310mm | 1
60 | Atraumatic Grip and Dissection Monopolar Forceps, Straight, Bounte, Serrated, Rotary, Insulated, Double Action, Complete Standard Instrument, Detachable, Composed of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert and Insert Sheath), Diameter 5mm, Length 310mm | 1
61 | Maryland Monopolar Pliers, Curved, Serrated, Rotary, Insulated, Double Action, Complete Standard Tool, Detachable, 4 PieceDistinct Modular Elements (Handle, Insulating Sheath, Insert and Insert Sheath), Diameter 5mm, Length 310mm | 1
62 | Babcock Monopolar Pliers, Fenestrated, Straight, Serrated, Rotary, Insulated, Double Action, Complete Standard Instrument, Detachable, Composed of 4 Distinct Modular Elements (Handle, Insulating Sheath, Sheath Insert and Insert), Diameter 5mm, Length 310mm | 1
63 | Nelson Monopolar Pliers, Right, Fine Serrated, Rotary, Insulated, Double Action, Complete Standard Tool, Detachable, Composed Of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert And Insert Sheath), Diameter 5mm, Length 310mm | 1
64 | Dorsey Intestinal Monopolar Forceps, Fenestrated, Straight, Serrated, Rotary, Insulated, Double Action, Complete Standard Instrument, Detachable, Composed Of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert And Insert Sheath), Diameter 5mm, Length 310mm | 1
65 | Croce-Olmi Monopolar Pliers, Double Fenestrated, Straight, Serate, Rotary, Insulated, Single Action, Complete Standard Instrument,Detachable, Composed Of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert And Insert Sheath), Diameter 5mm, Length 310mm | 1
66 | Grip Monopolar Pliers, 2x4 Teeth, Straight, Rotary, Insulated, Simple Action, Complete Standard Tool, Detachable, Composed Of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert And Insert Sheath), Diameter 5mm, Length 310mm | 1
67 | Grip Monopolar Pliers, Jaws With Longitudinal Grooves, Straight, Serrated, Rotary, Insulated, Double Action, Complete Standard Tool, Detachable, Composed Of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert And Insert Sheath), Diameter 5mm, Length 310mm | 1
68 | Mixter Monopolar Pliers, Angled 90°, Rotatable, Insulated, Double Action, Complete Standard Tool, Detachable, Composed Of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert And Insert Sheath), Diameter 5mm, Length 310mm | 1
69 | Grip Monopolar Pliers, With Reservoir in Jaws, Straight, Serrated, Rotary, Insulated, Double Action, Complete Standard Tool, Detachable, Composed of 4 Distinct Modular Elements (Handle, Sheath Insulators, Sheath Insert and Insert), Diameter 5mm, Length 310mm | 1
70 | Laparoscopic Port, Tungsten Inserts, Left Curved, Diameter 5mm, Length 310mm | 1
71 | Laparoscopic Port, With Automatic Positioning of the Needle in a Vertical Position, Tungsten Inserts, Straight, Diameter 5mm, Length 310mm | 1
72 | Metzenbaum Bipolar Scissors, Curved, Blunt/Bunt Tip, Rotary, Insulated, Double Action, Complete Standard Instrument, Detachable, Composed Of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert And Insert Sheath), Diameter 5mm, Length 310mm | 1
73 | Maryland Bipolar Forceps, Fenestrated, Curved, Rotary, Insulated, Double Action, Complete Standard Instrument, Detachable, Composed of 4 Distinct Modular Elements (Handle, Insulating Sheath, Insert Sheath and Insert), Diameter 5mm, Length 310mm | 1
74 | Bipolar Cable With 2 Pins, Aesculap Type Instrument Connection With 2 Blades, Unit Connection With 28.6mm Pin Distance, Compatible With Aesculap, Bowa, Lamidey, Soring, Valleylab Units, Olympus, Length 4m | 1
75 | Monopolar Cable, Instrument Connection 4mm, Unit Connection 5/2mm, Compatible With Aesculap, Acmi, Berchtold, Martin Units, Length 3.5m | 2
76 | Transfascial Suture Tool, For Subcutaneous Closure of Trocar Incisions, Diameter 2.5mm, Length 170mm | 1
77 | Irrigation/Suction Cannula, 2 Way Faucet, Diameter 5mm, Length 330mm | 1
78 | Fan Type Endoscopic Retractor With 5 Blades, Diameter 10mm, Length 400mm | 1
79 | Knot Pusher, Diameter 5mm, Length 310mm | 1
80 | Endoscopic Applicator For Titanium Surgical Clips With Two Pairs Of Parallel Arms For Increased Radial Stability, Diamond Shape Internal Profile, Distal End Closure, Medium-Large Size, Green Color, Diameter 10mm, Length 310mm | 1
81 | Titanium Surgical Clips with Two Pairs of Parallel Arms for Increased Radial Stability, Internal Diamond Profile, Distal End Closure, Medium-Large, Size Open Clip: 7.8x8.9mm, Green Cartridge With 6 Clips | 15
82 | Complete Handle for Monopolar Laparoscopic Electrode, Diameter 5mm, Length 330mm | 1
83 | Hook Type Ceramic Electrode Tip, Monopolar, L-Shaped, Diameter | 1
5 mM
84 | Injection Cannula, Luer Lock, Diameter 5mm, Length 330mm | 1
85 | Exploration Sample, Right, Diameter 5mm, Length 400mm | 1
86 | Aluminum Container Body, Without Perforated Base, Standard Size 1/1: 592x274x187mm | 1
87 | Anodized Aluminum Container Lid, Standard Size 1/1: 593x294x37mm, 2 Filters For 5000 Sterilizations Included, Red Color | 1
88 | Support For Fixing Minimally Invasive Instrumentation/Exchange In Basket Container, In View Of Sterilization, Standard Size 1/1: 540x253x166mm | 1
89 | Support For Fixing Exchange In Basket Container, In View Of Sterilization | 1
90 | Container Identification Plate, Aluminum, Red Color, With Inscription, Inscribed Up to 13 Characters/Line, 1-2 Lines, With 5mm Character Height, Size 56x18mm | 2
91 | Surgical Instrument Lubricating Solution/Oil, Capacity 300ml | 5
92 | After sales support

## Table 2
TECHNICAL SHEET No.: | YES | NOT | Remark
Containers for sterilization:
To meet the quality standards of the European Union-EU (DIN ISO 9001/9002).
To have a specific and distinct CE mark
The containers must comply exactly with the specified dimensions and characteristics
The containers must be legibly marked by the manufacturer with the following elements: the marking of the manufacturing plant (emblem or logo) and the marking of the manufacturer's name
To be provided with a cover very resistant to wear, shocks or deformations, thermostable. For easy identification, there should also be the possibility of using caps of at least 5 different colors
To have a double lid locking system, to ensure the perfect sealing of the inside of the container and at the same time the hygienic handling of the container after sterilization - during transport.
The lid of the box should be made of aluminum and be provided with filters made of reusable micropore material made of PTFE (teflon filter) for a long period of use so that after 5000 sterilizations, the tests performed by the manufacturer on the teflon filters should not show any changes significant resistance to flow through the membrane (steam penetration). Its filtersbe protected and covered by a steel plate firmly attached to the outer surface of the cover.
The germ barrier system of the PTFE filter case, according to EN868-8, Annex G, to be validated at 5000 sterilization cycles. The casoleta must be microbiologically tested for tightness, by the manufacturing company, (according to international standards) at 5000 sterilization cycles in aerosol spray with 108 KBE/ml germs. The supplier must prove these validations through the tests and studies provided by the manufacturer.
The sterility of the products in the containers must be guaranteed for at least 6 months, documentation provided by the manufacturer through the tests and studies done.
The manufacturer must carry out resistance tests to pressures higher than the standard gradients (according to DIN EN 285 – 10 bar/min). The casings must withstand three times higher gradients, 30 bar/min, without showing deformations. The supplier must prove these validations through the tests and studies provided by the manufacturer.
The boxes must be tested by the manufacturer, in accordance with DIN EN 868-8 B, so that they do not show mechanical deformations under a load of 40 kg for 2 minutes. The supplier must provide proof these validations through the tests and studies provided by the manufacturer.
The internal organization of the case can be done using baskets, packing fields and fixing systems specific to each type of instrument
To have the possibility of being provided with aluminum identification labels that can be written with at least 13 characters, of different colors.
The manufacturing company should offer repair services and bringing the containers up to nominal parameters, with a guarantee of at least 1 year.
The manufacturing company should offer in electronic versions, online, updated information about the maintenance and reprocessing of containers, in accordance with DIN EN ISO 17664.
The containers must comply exactly with the dimensions and characteristics specified in the presentation catalog.
To have the possibility of being provided with special labels for the triple-role box: sealing the box, identification of the box and indicator for sterilization
C. WARRANTY AND SERVICE CONDITIONS
- period: 2 years
- maximum intervention time: 24 hours at the beneficiary's premises
- the company that provides service during the warranty and post-warranty period to have a system of quality management according to EN ISO 9001: 2000, certificate
POST WARRANTY SERVICE
- minimum period: 10 years
- maximum intervention time: 48 hours, at the beneficiary's premises
- the local representative of the manufacturer who provides service during the warranty period and post guarantee through specialists trained and authorized by the manufacturer, to have implemented a quality management system according to EN ISO 9001: 2000, certified.
INSTRUCTION
- medical personnel